## Slide 1: PMP CPC Round ROBIN
Additional evaluations at AVL

## Slide 2: CONTENT
Charging efficiencies
Effect of calibration material
Instrumentation effects (DMA, electrometers)
10 nm calibrations with miniCAST

## Slide 3: Calibration setup
Electrospray (Emery Oil / Dioctyl Sebacate), RR APG, miniCAST (evaporating tube or catalytic stripper), PALAS DNP3000
Tandem DMA  Charging efficiencies
Two different types of electrometers were also available

## Slide 4: Charging efficiencies
The tandem DMA setup allows for the quantification of double to single charge fractions
  Charging efficiencies differ for soot/graphite and oil droplets
  ISO15900 recommendation was found to lie in between
  Soft X-ray neutralizers yield systematically higher fractions of multiply charges
85Kr
Soot
Mamakos, Aerosol Science & Technology 50:363-372
ISO15900

## Slide 5: Charging Efficiencies Effect
Default settings of RR APG led to rather high double charge fractions above 60 nm (>10% of ISO15900 requirement) for calibrations against electrometer
The use of default charging efficiencies was calculated to result in ~ 2% underestimation at sizes above 60 nm
85Kr

## Slide 6: Calibration material effect
Response for soot and graphite particles distinctly different than for emery oil
Efficiency for dioctyl sebacate much higher than any activation theory suggests  Chemical effect

## Slide 7: Instrumentation Effects - DMA
Repetition of emery oil tests with different DMA/classifier yielded systematic differences
Effect is more prominent at 23 nm, where the counting efficiency curve is steep

## Slide 8: Instrumentation Effects - Electrometer
We received a different type of electrometer during the campaign
A systematic difference of 8% was observed, which was rather independent of size even if flows were properly controlled

## Slide 9: 10 nm measurements
Setup:
  miniCAST burner settings to produced size distribution peaking below 10 nm
  Thermal treatment with Catalytic Stripper to avoid additional dilution
  Concentrations at 10 nm registered by the electrometer were ~15000 #/cm3
  Not optimal settings due to setup (i.e. use of long DMA) and time limitations

## Slide 10: CONclusion
Care is required in dealing with multiple charge fractions especially when an electrometer is employed as a reference
Importance of traceable calibration of reference instrumentation: DMAs & Electrometers
Organic liquid particles show distinctly higher counting efficiency, and there is a clear strong chemistry effect  Chemical composition needs to be defined in conjunction with the type of working fluid employed